Teensy Logic Analyzer User's Guide
8/8/2018

There are 2 modes to run in:
1. Basic (default) – hey, I just want to record up to 8 signals, up to about 1 MHz
2. Hardware – I only have 1 or 2 signals, and want to sample at up to 24 MHz (LC), 72 MHz (3.2), or 120 MHz (3.6) with full simple triggering capability

For Hardware mode capabilities, see tables at end of document.

Quick Setup

1a. Windows - Download OLS .zip (https://github.com/LAtimes2/ols/releases/latest), unzip, and run run.bat

1b. Linux - Download OLS .tar.gz (https://github.com/LAtimes2/ols/releases/latest), extract, and run run.sh

2. Select Capture -> Begin Capture, and set Analyzer port to COM port or /dev/tty for the Teensy

3. Select Device type for your Teensy type and capabilities (Basic/Demo/Hardware), then select Load Firmware

Basic mode:

Channel	Pin
0	2
1	14
2	7
3	8
4	6
5	20
6	21
7	5

Hardware mode:

Channel	Pin
0	1
1	8

OLS user interface (PC client)

The Teensy communicates with the PC via a SUMP application. The SUMP protocol is a standard for sending logic analyzer data. This document uses the OpenBench Logic Sniffer PC client for these examples. It runs under Windows/Linux/OSX and is available for download at github.com/LAtimes2/ols/releases or www.lxtreme.nl/ols. See installation.md for instructions on installing all applications.

Setting up OLS

OLS will remember the settings between runs. But the first time you must set up a few items.

Initial Screen
Select Capture → Begin capture:

Select Analyzer port if necessary:

First time - select 'Show device metadata' (this resets all parameters, so only use it if you are not sure of the device type, and check Acquisition and Triggers tabs after selecting).

Device type lists various configurations that the Teensy can be built for. Default for Teensy 3.1/3.2 is Teensy96, and default for Teensy LC is Teensy48.

Firmware is the version of the Teensy software.

Protocol is the SUMP protocol and is always 2.

If you have not loaded firmware in your Teensy yet, you can do it by selecting the appropriate Device type, then select Load Firmware. It will display a dialog like this (it will automatically close when complete):

Select the Acquisition tab.

For the basic Teensy 3.2 configuration, the defaults are Sampling Rate of 2 MHz and Recording Size of 58k samples.

Number scheme - not used

Sampling Clock - not used, always internal

Sampling Rate - set as desired. Higher speeds have some limitations, see tables at end of file.

Channel Groups - keep 0 selected

Number of Channels - Fewer channels allows recording more data. Selecting 4 channels can record twice as much data as 8 channels, 2 channels is 4x, 1 channel is 8x more data.

Recording Size - Default value is the maximum for 8 channels. If you set it higher, you also need to reduce the number of channels accordingly. It will give you a message dialog if trying to record too much data for the number of channels.

Run Length Encoding - if selected, it compresses the data to record for longer times. If your data isn't changing very often, Run Length Encoding (RLE) can record up to 100 times longer when selected. It does not take affect at the 3 highest sample rates since it takes more computing power to record.

RLE uses the highest channel for internal data, so it can only record 7 channels. RLE cannot be used when recording 1, 2, or 4 channels.

Select Capture to start recording.

No inputs:

If you are using Demo firmware, you will see capture data like this:

Zoomed in:

Time measurement – drag right border to the left:

Trigger setup:

This example sets up a trigger on channel 4 being high at 10% into the buffer.

Capture looks like this:

While the logic analyzer is waiting for a trigger, the red Stop button is enabled. To halt the analyzer when no trigger is received, press the button. No data will be sent back from the Teensy unless Run Length Encoding (RLE) is selected, since RLE can go a long time if the data is not changing.

Advanced Trigger Setup:

Some modes allow complex trigger sequencing (up to 4 stages). For example, if a signal is usually high, but you want the low to high transition, you can set up 2 stages, with the first stage triggering on the signal being low, and the second stage triggering on high. There is also a delay between stages, so if you want the fourth clock select on a SPI transfer, calculate the number of samples in each transfer and put that delay (plus a little extra) between stages.

To set up a stage,select COMPLEX type, then select Stage X tab. Set trigger mask and value for each stage, along with any optional delay after the stage is triggered. On the last stage, check the Start Capture box to indicate the trigger is complete.

LED

The LED on the Teensy works as follows:

Blinking every 2 seconds – idle (not recording)
On – recording, waiting for a trigger
Quick flashes – receiving commands from the PC client.
Basic Configurations

There is also a 144 MHz configuration file. You need to compile for 144 MHz (optimized) to use it (if you don't know how, google it, just use 96 MHz, or use the pre-built Hex files).

Not all combinations of speed and samples are valid. See tables below. When a combination cannot be met, the display will have alternating 1's and 0's in the invalid sections (an entire channel, or a portion of time, or both).

Teensy 3.2 96 MHz	2 MHz and below / Complex Trigger	12 MHz / Simple Trigger	19.2 MHz / Pre-Trigger Only
58k samples and below	8 channels	8 channels (Note 1)	8 channels (Notes 1, 3)
116k samples	4 channels	Note 2	Note 2
232k samples	2 channels	Note 2	Note 2
464k samples	1 channel	Note 2	Note 2

Teensy 3.2 144 MHz	3 MHz and below / Complex Trigger	18 MHz / Simple Trigger	28.8 MHz / Pre-Trigger Only
58k samples and below	8 channels	8 channels (Note 1)	8 channels (Notes 1, 3)
116k samples	4 channels	Note 2	Note 2
232k samples	2 channels	Note 2	Note 2
464k samples	1 channel	Note 2	Note 2

Note 1: Timing data is valid up to about 28k samples, then is 10% too fast (upper SRAM?)
Note 2: Limited to 58k samples. No advantage over 58k setting.
Note 3: Trigger Before/After Ratio is not used. Trigger, if selected, is always at the beginning.

Teensy 3.5 120 MHz	2.5 MHz and below / Complex Trigger	15 MHz / Simple Trigger	24 MHz / Pre-Trigger Only
186k samples and below	8 channels	8 channels	8 channels (Note 1)
372k samples	4 channels	Note 2	Note 2
744k samples	2 channels	Note 2	Note 2
1488k samples	1 channel	Note 2	Note 2

Note 1: Trigger Before/After Ratio is not used. Trigger, if selected, is always at the beginning.
Note 2: Limited to 186k samples. No advantage for higher settings.

Teensy 3.6 240 MHz	5 MHz and below / Complex Trigger	30 MHz / Simple Trigger	48 MHz / Pre-Trigger Only
250k samples and below	8 channels	8 channels	8 channels (Note 1)
500k samples	4 channels	Note 2	Note 2
1000k samples	2 channels	Note 2	Note 2
2000k samples	1 channel	Note 2	Note 2

Note 1: Trigger Before/After Ratio is not used. Trigger, if selected, is always at the beginning.
Note 2: Limited to 250k samples. No advantage for higher settings.

Teensy LC	1 MHz and below / Complex Trigger	8 MHz / Pre-Trigger Only
4k samples and below	8 channels	8 channels (Note 1)
8k samples	4 channels	Note 2
16k samples	2 channels	Note 2
32k samples	1 channel	Note 2

Note 1: Trigger Before/After Ratio is not used. Trigger, if selected, is always at the beginning.
Note 2: Limited to 4k samples. No advantage for higher settings.

Teensy 3.0 96 MHz	2 MHz and below / Complex Trigger	12 MHz / Simple Trigger	19.2 MHz / Pre-Trigger Only
10k samples and below	8 channels	8 channels (Note 1)	8 channels (Notes 1, 3)
20k samples	4 channels	Note 2	Note 2
40k samples	2 channels	Note 2	Note 2
80k samples	1 channel	Note 2	Note 2

Note 1: Timing data is valid up to about 5k samples, then is 10% too fast (upper SRAM?)
Note 2: Limited to 10k samples. No advantage over 10k setting.
Note 3: Trigger Before/After Ratio is not used. Trigger, if selected, is always at the beginning.

Hardware Configuration

Once you are comfortable with the basic configuration, you can try the hardware configuration to get higher speeds and/or more samples. To use it, compile with

 #define HARDWARE_CONFIGURATION 1

and select Device type as Teensy 96 MHz OLS Hardware or Teensy LC 48 MHz OLS Hardware.

There is also a 120 MHz Hardware configuration file. It is the same as below for Teensy 3.1/3.2, but 25% faster.

At some higher speeds and/or sample sizes, only 1 channel is available. See tables below. When a combination cannot be met, the display will have alternating 1's and 0's in the invalid sections (an entire channel, or a portion of time, or both).

Teensy 3.2 96 MHz	48 MHz and below / Simple Trigger
232k samples and below	2 channels
464k samples	1 channel

Teensy 3.2 120 MHz	60 MHz and below / Simple Trigger
232k samples and below	2 channels
464k samples	1 channel

Teensy 3.2 144 MHz	72 MHz and below / Simple Trigger
232k samples and below	2 channels
464k samples	1 channel

Teensy 3.5 120 MHz	60 MHz and below / Simple Trigger
744k samples and below	2 channels
1488k samples	1 channel

Teensy 3.6 240 MHz	120 MHz and below / Simple Trigger
1000k samples and below	2 channels
2000k samples	1 channel

LC: Gets you 12 MHz, 2 channels at 16k samples (vs 4k). Pretrigger.
Gets you 6,4,2 MHz, 2 channels at 16k samples with full trigger.
Gets you 24 MHz 1 channel. Pretrigger

Teensy LC	6 MHz and below / Simple Trigger	12 MHz / Pre-Trigger Only	24 MHz / Pre-Trigger Only
16k samples and below	2 channels	2 channels (Note 1)	1 channel (Note 1)
32k samples	1 channel	1 channel	1 channel (Note 1)

Note 1: Trigger Before/After Ratio is not used. Trigger, if selected, is always at the beginning.